| Laboratory: | Document Type: | Original Date Adopted: | Previous Document: |
| --- | --- | --- | --- |
| Point of Care Testing | Procedure | 07/31/2007 | POC-1 Revision 7 |

| Document Author: | Document Owner: | Acknowledgement / Required Copy Holders*: |
| --- | --- | --- |
| Alicia Sheffield, MLS(ASCP)cm | Karen Scott, MT(ASCP) | Testing personnel documents acknowledge of reading procedures during training and annual competency |

Approval*:

Point of Care Coagulation Division Director University Hospitals Laboratory Medical Director University Hospitals East Laboratory Medical Director

*Approval and Acknowledgements*

Refer to QPulse system and Document Details report for laboratory directors(s)’ electronic signature approval, employee acknowledgment and effective date.

- PRINCIPLE:

  - HEMOCHRON Microcoagulation Systems utilize a mechanical endpoint clotting mechanism in which testing occurs within the disposable ACT-LR or ACT+ cuvette. Following whole blood sample introduction, the instrument precisely measures 15 microliters of blood and automatically moves it into the test channel within the cuvette. The remainder of the blood sample, not needed for testing, is automatically drawn into the waste channel of the cuvette.

  - Sample/reagent mixing and test initiation are performed automatically, requiring no operator interaction. After mixing with the reagent, the sample is moved back and forth within the test channel and monitored by the analyzer for clot formation. The clot detection mechanism consists of two LED optical detectors aligned with the test channel of the cuvette. The speed at which the blood sample moves between the two detectors is measured. As clot formation begins, blood flow is impeded and the movement slows. The instrument recognizes that a clot endpoint has been achieved when the movement decreases below a predetermined rate. Electronic optical detection of a fibrin clot in the blood sample automatically terminates the test.  The instrument’s digital timer displays the Celite® equivalent ACT value in seconds in order to provide a familiar clinical format and thus facilitate accurate clinical test result interpretation.

  - The ACT+ test is intended for use in monitoring patients receiving heparin anticoagulation therapy. ACT+ demonstrates a dose response to heparin concentration of more than 1.0 unit of heparin/ml of blood but less than 6.0 units. The ACT + test should be used when the ACT result needs to be above 400 seconds.

  - The ACT-LR demonstrates a more sensitive dose response to heparin concentrations up to 2.5 units; range is between 65 and 400 seconds.

- Scope of Document:

  - This document applies to all personnel performing the ACT+ and ACT-LR testing and testing oversight.

- RESPONSIBILITY:

  - The POCT coordinators and manager are responsible for maintaining this document and ensuring biennial review. The coagulation laboratory division director is responsible for approving all changes, and reviewing at least biennial. The laboratory medical director is responsible for establishing and approving all changes before activating document.

- SPECIMEN COLLECTION:

  - All institutional policies and procedures should be followed in the collection of blood samples including the presence of an order.

  - Patient Preparation: Verify patient identification using at least 2 identifiers.

    - During surgical procedures, patients are identified per OSU Time Out.

  - Specimen Type:  Whole blood from arterial or venous access lines or sample ports of the extracorporeal circuit.

  - Obtain at least 5 to 6 ml of dead space volume of blood and discard prior to collection of the test sample in order to eliminate the risk of excess dilution and contamination of the sample with heparin from the catheter or line.

  - Handling conditions: Test fresh whole blood specimens immediately following sample collection.

  - Specimen Amount for testing

    - Withdraw at least 0.2 ml of blood from designated location.

    - The cuvette requires a minimum volume of 15 microliters of fresh whole blood to perform the analysis, and will display “Sample too small” if insufficient sample is applied

  - Unacceptable Specimens: clotted specimens or specimens collected in a heparinized ABG syringe cannot be analyzed and requires recollection.

- REAGENTS/SUPPLIES

  - Room temperature (15-30°C) is monitored each day of use by the testing department and recorded on temperature log.  The temperature log is reviewed monthly by the POC department.

    - If a minimum/maximum thermometer is used to perform continuous monitoring of temperature between daily temperature readings or following a laboratory downtime (e.g. laboratory closure for weekend or holiday), both the low and high temperatures must be recorded.  To ensure correct temperature readings, the minimum/maximum thermometer device must be reset prior to the monitoring period.

  - ITC Hemochron Signature Elite

    - Must be used at room temperature 15-30°C (59-86°F)

    - There are no humidity requirements to use the Hemochron Signature Elite

    - If the HEMOCHRON test system becomes inoperable, do not use the system and contact the Point-of-Care department at 614-685-6610 immediately.  A loaner device may be available for use.

  - Hemochron Jr. Plus activated clotting time (ACT+) Cartridges

    - ACT+ cartridges are a self contained disposable test chambers preloaded with a mixture of silica, kaolin and phospholipids as an activator.

    - Each lot of cuvettes should be validated for performance at two liquid quality control levels when a   new shipment is received and once per 30 calendar days thereafter.  Acceptable performance ranges for the cuvettes are included with each quality control product kit.

    - Cuvettes can be stored at 2-8oC up to, but not beyond, manufacturer expiration date. The temperature is monitored by Isensix in the point of care office.

    - Upon removal from refrigerator, the cuvettes are stored at room temperature (15-30 oC).  Room temperature re-dating is to a maximum of 12 weeks, but must never exceed the marked manufacturer expiration date.

  - Hemochron Jr.™ Low range heparin activated clotting time (ACT LR) Cartridges

    - ACT LR cartridges are a self contained disposable test chambers preloaded with dried preparation of Celite and potato dextrin as an activator.

    - Each lot of cuvettes should be validated for performance at two liquid quality control levels when a   new shipment is received and once per 30 calendar days thereafter.  Acceptable performance ranges for the cuvettes are included with each quality control product kit.

    - Cuvettes may be stored at 2-8oC up to, but not beyond, manufacturer expiration date. The temperature is monitored by Isensix

    - Upon removal from refrigerator, the cuvettes are stored at room temperature (15-30 oC).  Room temperature re-dating is to a maximum of 12 weeks, but must never exceed the marked manufacturer expiration date.

  - DirectCheck ™ liquid controls

    - High-range heparin controls: Level 1 (DCJACT-N) and Level 2 (DCJACT-A) for ACT+ cartridges

    - Low-range heparin controls: Level 1 (DCJLR-N) and Level 2 (DCJLR-A) for ACT-LR cartridges

    - The Direct Check liquid controls consist of dried fixed bovine red blood cells, buffered sheep and horse plasma. The diluent portion in the dropper vial contains distilled water, sodium chloride, Tween 20, anticoagulant and calcium chloride. The diluent rehydrates the dried blood once the vial is cracked.

    - Clot time varies with each lot number of ACT + and ACT-LR control. The clot times are included in the control package insert.

    - When refrigerated (2-8°C) the controls are stable until the marked expiration date.  Vials may also be stored at room temperature (15-30°C) for up to 4 weeks, but must never exceed the marked expiration date. The POC refrigerator is monitored by Isensix.

- SPECIAL SAFETY PRECAUTIONS:

WARNING: BODY FLUID PRECAUTION

  - Blood is a body fluid capable of transmitting infectious diseases.  Universal precautions for the prevention of the transmission of blood borne pathogens must be in effect at all times.

  - Personal protective equipment: wear gloves throughout the specimen collection and testing process

- Calibration/programming/cleaning (POC testing personnel)

  - Calibration: The instrument is self-calibrated, as all instrument functions are continuously monitored and verified by the instrument software when a test is performed.  The instrument does not require additional calibration by the user.

  - Correlations: Performed biannually.  Refer to POC Quality Management Plan.

        - Quality Cross Check samples from CAP are used: CT2Q—ACT-LR, CT3Q—ACT+

        - Insert ACT cuvettes into each of the HEMOCHRON® Jr. analyzers.

        - Add Quality Cross Check sample into sample well of a pre-warmed ACT cuvette, which has been previously inserted into a HEMOCHRON® Jr. analyzer (Display should read: “ADD SAMPLE…PRESS START”).  Follow CAP provided sample preparation instructions.

        - Immediately after the addition of the blood into the cuvette well, press the “START” button. Record the ACT result after test is complete.

        - Record the clotting times and calculate the r value in an appropriate spreadsheet.

        - Give the report to the coagulation division director for his/her review and signature

        - Follow correlations criteria for outliers, instrument issues and provide summary report.

  - Cleaning  procedure:

    - Inspect and clean the cuvette opening at least monthly.

    - If disinfection is needed, disinfect with a 0.5% solutions of sodium hypochlorite (Orange Sani Wipes)

    - DO NOT submerse the unit in any liquids.

  - System Self-Checks:

    - Each ITC Hemochron Signature Elite is equipped to perform a self-check every time the instrument is activated and a test is performed.  When a test is initiated by insertion of a cuvette, the system automatically checks:

      - Verification of adequate battery power to complete a full test

      - Verification of the test type on the screen display to ensure that the Liquid Electronic Diodes (LED’s) used for identifying the tests are functioning properly.

      - Verification that the cuvette temperature is warmed to 37°C ±1°C. If this temperature is not achieved or is exceeded, an appropriate error message will be displayed and testing is prohibited.

      - Verification that the internal timers function correctly for each test. If the system timer and assay timer disagree, a real-time clock error message is displayed and the test result is not reported.

      - Verification that the sample is present and is of sufficient size to run the test. This ensures that the pumps and sample sensing LED’s are functioning properly and that the cuvette is adequately sealed.  If these instrument and sample parameters are not appropriate, the test is terminated and an error message is displayed.

  - Performance verification after a repair or on a new instrument

    - Perform a precision study which includes measuring a combined total of 20 ACT measurements, a combination of EQC and Liquid QC, as follows:

      - Measure 10 EQC measurement—validates the low and high ACT timing measurement ranges with internal electronic QC processes (all must PASS).

      - Measure 5 Liquid QC Level 1—validates the low ACT timing measurement ranges with external liquid QC processes (results must be within Package Insert Limits).

      - Measure 5 Liquid QC Level 2—validates the high ACT timing measurement ranges with external liquid QC processes (results must be within Package Insert Limits).

    - Liquid Quality Controls for the instrument will be documented and assessed for acceptability within the manufactures range.

    - If available loaner meters may be given to the department while their meter is out for repair.

      - If a loaner is put in use, QC must be performed prior to patient testing.

  - Start Up/Shut Down

      - To start up or shut down the Signature Elite, press and hold the START button.

- QUALITY CONTROL:

  - Internal Electronic Control: Automatically performed every 8 hours of use by the testing department/unit

    - Turn on the instrument by pressing and holding the START button.  EQC will automatically start if it is due.

    - To manually run EQC

      - Select QC from the Main Menu.

      - Select option #1 for EQC.

    - Instrument automatically runs a system check and both low and high electronic controls.

    - Electronic QC indicates “PASS” if the normal level (30-second) and the abnormal level ( 300-second or 500-second) plus the internal temperature check meet manufacturer’s acceptable performance criteria.

    - If the meter fails EQC check, repeat the EQC.  If the meter fails a second time, do not use the meter.  Contact the Point-of-Care testing staff at Path.Glucose@osumc.edu or call the POC office at 685-6610.

  - External Liquid Controls (LQC):

    - DirectCheck ™ Liquid Controls

      - ACT controls: Level 1 and Level 2 are performed:

        - On each analyzer once a month by testing personnel.

        - On new cuvette lots, or new shipments of same lot of cuvettes.

        - To confirm acceptable range of new lots of QC before use.

Note: The control testing process is very technique and time dependant. Review process before starting

    - Liquid QC Procedure (Testing Personnel)

      - Remove controls from the refrigerator and bring to room temperature (this could take up to 60 minutes).  Performing testing on cold QC may not pass acceptability.

      - Visually inspect the QC material to ensure that the ampule is intact. Place the appropriate cartridge (ACT+ or ACT-LR) into the cartridge slot of the Hemochron analyzer. Scan the barcode on the cuvette package to download the lot number and expiration date of the cuvette.

      - Enter badge ID number by manually typing on the keypad or scanning the barcode on the ID badge.  If the ID is manually entered you must press and hold ENTER button.

      - Press the QC button at the bottom of the instrument so the instrument goes into QC mode.

      - Select option #1 for normal QC or option #2 for abnormal QC.

      - Select the correct lot number or scan the appropriate QC barcode to download the lot number and expiration of the control being used.

      - Remove the outer plastic seal from the ampule of the control material and place in a protective sleeve provided with each box of controls.

      - Holding the vial upright, gently tap the ampule to allow the liquid contents to settle to the bottom of the vial.

      - Crush the inner glass ampule three times to ensure complete breakage of the ampule.

      - Shake the control vigorously to mix until the red blood cells are uniformly dispersed and the control is completely reconstituted, approximately 10 seconds.

      - While inverting the vial (dropper tip down), use a downward snapping motion of the wrist to ensure control material flows to the tip of the dropper.

      - Remove the cap.

      - Discard the first drop, then add one drop of control material to the sample well of the cartridge.

      - Fill the cuvette sample well to the top. Should a large dome extend over the top of the center sample well, push it over into the outer sample well.

      - Immediately press the START key.

      - Recap the control vial and remove the control vial from the protective sleeve.

      - Dispose of the vial into biohazardous waste.

      - Wait for a single beep signaling the conclusion of the test.

      - Results are displayed as the Celite equivalent clotting time in seconds.

      - Note the results to ensure that the controls have passed and are within acceptable range. The control ranges are lot specific and vary from lot-to-lot. The range data is on the package insert used to scan the lot number.

      - Both levels must be acceptable prior to reporting patient results.

      - Download the Quality Control data to Telcor from individual systems each time QC is performed by attaching the ether net cable to the system network.  Select database and press the #6 (POCT>>NET).  The data transfer will begin and will download into Telcor

      - When QC is unacceptable the instrument will lock out.

    - Acceptance limits and how confirmed (POCT Testing Department)

      - Liquid control ranges are confirmed per shipment of quality control. Each box of control is barcoded to input the lot number and expected ranges to establish control limits.

      - 5 vials of QC of each level are run to verify new lot of QC compared to manufacturer’s ranges.  Record results in Sig Elite QC Lot Verification Worksheet.

    - Quality Control Results Review

      - Liquid Quality Controls for each system will be documented.

      - The QC data is reviewed and acceptable prior to reporting patient results.

    - Corrective actions to be taken for unacceptable QC

All steps completed must be documented in QML with the action taken.

      - If you get a fault message, the sample had bubbles.

      - If the control result is below the acceptable range:

        - The reconstituted control was not thoroughly mixed.

        - The time period between breaking ampule of the control, material mixing, sampling, and pressing start.

      - If the control result is above the acceptable range:

        - The inner glass ampule was not adequately crushed.

        - Reconstituted control was not thoroughly mixed.

        - Vial cap is removed prior to inverting, allowing diluent to leak from the vial.

      - If unacceptable QC is due to processing error and repeat QC is acceptable, patient results are acceptable.

      - If the instrument continues to fail either level of liquid QC do not use the instrument

      - Corrective action is documented in QML.

    - The QC is reviewed by the manager and the coagulation division director. The results are maintained on the L: Drive.

- TEST PROCEDURE:

  - Ensure you have the correct cartridge for the instrument and test.

  - The instrument will only run the test cartridge it is programed to run.

    - Example: Do not try to run an ACT + when your instrument is programmed for ACT-LR

  - Press and hold START button to turn the instrument on. The Hemochron Signature Elite will complete a system check

    - To perform a high-range ACT patient test, use the ACT+ cartridge.

    - To perform a low range patient test, use the ACT-LR cartridge.

    - Insert appropriate cartridge into the Hemochron Signature Elite. During the pre-warm stage, observe the display for any fault messages.

    - Scan the barcode on the cuvette package to download the lot and expiration date of the cuvette.

    - Scan your user ID.

    - Scan the Patient ID barcoded wristband of the patient.

    - Note: all required quality control tests must be performed before patient testing

    - The instrument will signal when ready with an audible tone, and the screen will display the messages “Add Sample” and “Press Start”.  The instrument will remain in the ready mode for five minutes. At the end of five minutes, a “Start Timeout” will occur indicating that the current cartridge must be discarded and a new cartridge placed in the instrument

    - NOTE: Blood samples must not be collected until the instrument display indicates “Add Sample” And “Press Start”

    - Check for clots by expelling a small amount of blood on an absorbent surface.  Clotting begins immediately once the sample is drawn. Testing is highly sensitive to timing technique.

    - IMMEDIATELY dispense one drop of blood into the sample well of the prewarmed ACT cuvette.  Should a large drop of blood extend above the top of the center sample well creating a dome, simply push it over into the outer sample well.

      - When transferring blood into the sample well, DO NOT force blood into the hole located on the center of the sample well.  DO NOT generate air bubbles in the sample well.

    - IMMEDIATELY press the START key

    - Test completion will be indicated by a single beep

    - ACT results are automatically converted to a reference Celite ACT result and displayed as the Celite equivalent results in seconds which is transmitted into the electronic medical record (EMR) via Telcor/Sunquest interface after downloading the instrument.

      - Attach the Ethernet cable into the Ethernet port.

      - Press the Data Base button at the bottom of the instrument.

      - Press the #6 (POCT>>NET).

- CALCULATIONS:

  - No calculations are required. All results are automatically converted to a reference Celite® ACT value.

- REPORTING RESULTS:

  - Record results:

    - Record results in the Electronic Medical Record flow sheet or download into Telcor.

    - Procedures for reporting abnormal results:

      - Notify the physician performing the procedure. Document in the Medical Record.

    - Reference Range / Technical Range / Reportable Range: Refer to Laboratory Master List.

  - In case of a system downtime the results will be stored in the Signature Elite and will transmit to the medical record once the system is up and running and the instrument is downloaded.  Alternately, the testing personnel should document the result in the Electronic Medical Record flow sheet.

- INTERPRETATION OF RESULTS

  - All test results should be scrutinized in light of a specific patient’s condition, procedure and anticoagulant therapy. Any results exhibiting inconsistency with the patient’s clinical status should be repeated or supplemented with additional test data.

- LIMITATION OF PROCEDURE:

  - The ACT+ test is intended for use in monitoring patients receiving heparin anticoagulation therapy with a heparin concentration of more than 1.0 unit of heparin/ml of blood but less than 6.0 units.

  - Celite equivalent ACT values exceeding 1005 seconds are not reported on the instrument. Instead, an “Out of Range-Hi” message will be displayed on the Hemochron Jr. Signature Elite.

  - Samples with hematocrits less than 20% or greater than 55% are not recommended due to an optical density outside the level of detection of the instruments.

- REFERENCES:

  - ITC Hemochron Signature Elite™ Operator’s Manual, 2004

  - ITC Hemochron Signature Elite™ Whole Blood Microcoagulation System Configuration Manager Software, 2005.

  - Product Inserts for the following ITC Hemochron Signature Elite™ Cartridges and Controls:

    - ITC Hemochron Activated Clotting Time Plus (ACT+) Cartridges

    - ITC Hemochron Low Range Activated Clotting Time (ACTLR) Cartridges

    - DirectCheck Whole Blood Controls for Hemochron Jr Microcoagulation Systems, 2006

  - Bull BS, Korpman RA, Huse WM, Briggs BD; Heparin Therapy During Extracorporeal circulation: I. Problems Inherent in Existing Protocols. J Thorac Cardiovasc Surg 1975; 69:674-684.

  - Doty DB, Knott HW, Hoyt JL, Koepke JA: Heparin dose for accurate anticoagulation in cardiac surgery. J Cardiovasc Surg 20:597-604, 1979.

  - Esposito RA, Culliford AT, Colvin SB, Thomas SF, Lackner II, Spencer FE: The role of the activated clotting time in heparin administration and neutralization for cardiopulmonary bypass. J Thorac Cardiovasc Surg 85: 174-175, 1983.

  - Gambino R: Monitoring heparin therapy. Lab Report for Physicians 4:17-20, 1982.

  - Hattersley P: Activated Coagulation Time of Whole Blood. JAMA 1966: 136-436.

  - Wang J-S, Lin C-Y, Hung W-T, Thisted RA, Karp RB: In vitro effects of aprotinin on Activated Clotting Time measured with different activators. J Thorac Cardiovasc Surg 104, 4: 1135-1140, 1992.

- RELATED DOCUMENTS

  - Refer to QPulse System or Document Detail Report for related Laboratory Policies, Procedures, and Master Forms.